## Slide 1
Diagnosis
Persuasive Master Challenge
Vesicle Isolation
Protein Analysis

## Slide 2
CONTEXT

## Slide 3: PERSUASIVE focuses on non-acute chest pain patients
non-acute (95%)
acute (5%)
50%
5%
CHEST PAIN
possible underlying problems
20%
10%
10%
5%
CONTEXT
acute cardiac problems
can identify cardiac cases in non-acute chest pain

## Slide 4: 80% of chest pain referrals are unnecessary; opportunity for significant cost reduction
Illustrated is the Dutch situation - similar inefficiencies exist worldwide
Problem
Solution
525.000
100%
80%
525.000
80%
20%
Significant cost reduction * / **
(€200 mln/yr)
€1.400 / referred
patient **
€460 / patient
SOLUTION PROBLEM
* High cost savings are possible as estimated by Zorginstituut Nederland (2017), Zinnige Zorg verbetersignalement 'Pijn op de borst’.
** Confirmation in PERSUASIVE’s health-economic study (NL context) – see appendix, slide 18
100%
210.000 referred
210.000 referred
GP
Hospital
GP
Hospital
laboratory

## Slide 5: New solution using proven technology
Sample Processing
Quantification
Laboratory Robot
(Hamilton)
Read-out Unit
(Protein Simple – Ella)
Data Analysis
Dedicated Algorithm
Patient
sample
input
Patient
test result =
diagnosis
System to be
CE-marked
TECHNOLOGY
MVP = CE-marked modular system, to be realized in 2 years
More detailed technology impression – see appendix, slide 15

## Slide 6: User Story
USER STORY
Klinisch laborant ontvangt bloed van de huisarts
Laborant verwerkt het bloed met de Persuasive kit
Eiwit waardes worden bepaald met behulp van apparatuur
Deel-input voor de software zijn eiwit waardes
Huisarts importeert eiwit waardes en voegt patiënten gegevens toe. Software berekend laag, gemiddeld of hoog risico voor Stabiele Angina
Huisarts haalt resultaten van eiwit waardes uit het patiënten dossier
1
2
3
4
8
5
7
Laborant plaats resultaten in het Laboratory Information Management System
Lab resultaten worden overgebracht naar het patiënten dossier
6

## Slide 7: Input-Output
INPUT-OUTPUT
Patient number
Valid/non valid with error explanation
Low/medium/high risk of SA
Patient number
Patient name?
Sex
Age
Signal value protein 1
Signal value protein 2
Signal value protein 3
Signal value protein 4
Signal value control > x
INPUT
OUTPUT
ALGORITHM

## Slide 8: SOFTWARE workflow example
ALGORITHM
Check fields for value
Check if control value > x
Check protein signals > x
Calculate risk score
Classify risk score
Risk score variables:
Sex
Age
Signal value protein 1
Signal value protein 2
Signal value protein 3
Signal value protein 4
Error
Error
Error

## Slide 9: ALGORITHM Requirements
ALGORITHM
All input fields must be filled in, otherwise message that fields are open
Control sample must be above cut off value (to be determined), otherwise message ''recommend to repeat test (possibly contamination)''
Protein Signal values must be above value x (to be determined), otherwise message ''recommend to repeat test (possibly not enough target input)''
Input values determine risk score via algorithm with weighting factors (see example in case study)
Outcome of algorithm above below a value (to be determined) means a low risk of Stabile Angina
Outcome of algorithm above between values (to be determined) means a medium risk of Stabile Angina
Outcome of algorithm above between values (to be determined) means a high risk of Stabile Angina

## Slide 10: General requirements
REQUIREMENTS
User friendly (attractive/professional user interface) and recognizable for general practitioners
Software will be an In vitro Diagnostic Device (Class B) and must be compliant to ISO62304
Ensure that algorithm IP is protected
Patient data protected and meets GDPR
Input of signal value proteins must be importable (excel)
Input of other patient data separately added
Options platform: for example locked excel or cloud-based solution
Option to collect patient data at Persuasive for further research

## Slide 11: Case study
Case Patiënt
Patient data
Murad de Vries
23-03-1953
Male
Signal value protein 1 = 2
Signal value protein 2 = 4
Signal value protein 3 = 3
Signal value protein 4 = 3
Control value control = 2
Risk-calculation factors
Male factor = 0,5
Female factor = 0,75
Age factor = 0,5
Protein 1 factor = 2
Protein 2 = 0,5
Protein 3 = 1
Protein 4 = 0,5
Valid control value = 1-3
Minimal protein value = 0,2
Risk category | Risk-value
Low risk | <30
Medium risk | 30-40
High risk | >40

## Slide 12: Expectations
Case Patiënt
Working (functional) software prototype that can be used for demos by Persuasive
Use excel file as data input format, other formats may be considered
Create folder structure in which relevant background information found is stored for future use of Persuasive

## Slide 13: Develop a (online) software platform for the Persuasive Diagnosis Tool 1) Market research comparable platforms/user requirements 2) Should the tool be cloud-based/local software/plug-in/phone app 3) What is the program code Persuasive can use? 4) What should the frontend look like? Good luck with the challenge!
CHALLANGE